## Slide 1
Methods
Conclusions
Measurement-Oriented Microscopy Reconstruction
for Cross-Magnification Micro-Displacement Inspection
Bibliography
Introduction
Results
Paper Code:
Yen Hsuan Chiang, Liang-Chia Chen*
Department of Mechanical Engineering, National Taiwan University
No. 1, Sec. 4, Roosevelt Rd., Taipei 10617, Taiwan. *Correspondence: lchen@ntu.edu.tw
Microscopy reconstruction supports optical inspection when clearer boundaries also produce more stable position measurements. Objective changes alter object-space sampling, optical blur and the pixel phase of motion. The variation of measured single steps provides a direct test of reconstruction for displacement inspection.
MOSR reconstructs each monochrome image on its native pixel grid. One fixed set of learned weights serves 20× and 50× objectives. This study examines each objective at a common nominal step of 1 µm, with Raw, fixed Wiener and MOSR sharing the same image pairs and circle-center calculation.
Raw
MOSR
20×
2 µm
2 µm
50×
2 µm
2 µm
Fig. 1. Raw and MOSR images from the 20× and 50×, 1 µm datasets. Each comparison uses the same field and grayscale range. Scale bars: 2 µm.
Reconstruction. MOSR combines U-Net context [1], anti-aliasing [2] and bounded PSF guidance. The optical inputs comprise 25 measured kernels at 20× and one calibrated kernel at 50×. Fixed Wiener uses a regularization value of 0.0565.
Localization. Coarse contours [3] initialize maximum-gradient boundary points. Geometric least-squares circle fitting [4] uses 64 points at 20× and 96 at 50×. Every pillar starts from the same fixed calculation state.
Displacement. The median of nine pillar-center changes defines each step along the acquisition motion direction. The 20× dataset contains 30 complete scans of 30 steps. The 50× dataset contains 30 locked adjacent-frame pairs.
TABLE I. SEPARATE ACQUISITION CONDITIONS
Objective | Step (µm) | Measured steps | µm/pixel
20× | 1.0 | 30 scans × 30 | 0.1725
50× | 1.0 | 30 pairs | 0.0690
Each scan or locked set yields the sample SD of 30 steps (denominator 29). At 20×, all 30 scan SDs are shown; the within-scan summary is the square root of their mean squared SD. The 50× result is reported separately. Both use image-measured displacement.
At a common 1 µm step, MOSR lowers displacement variation at both objectives. Results are calculated within each acquisition condition.
20×, 1 µm: 30 repeated scans
52.3%
10.200 nm Raw
4.865 nm MOSR
Reduction in pooled within-scan Step SD
Chart: Scan number
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Raw | 7.6738 | 7.4194 | 9.6533 | 9.0003 | 9.6654 | 7.6594 | 9.5061 | 9.8325 | 9.3184 | 9.929 | 7.3538 | 9.2465 | 10.8028 | 6.9523 | 8.9056 | 11.4572 | 10.9004 | 11.3685 | 10.2767 | 9.896 | 10.9015 | 11.1019 | 13.7825 | 12.2664 | 12.3918 | 10.3141 | 10.6441 | 10.598 | 9.4849 | 13.5482
Wiener | 7.1149 | 5.7396 | 7.0344 | 6.0359 | 5.6113 | 6.1504 | 5.7944 | 5.6622 | 8.1385 | 5.9734 | 4.9032 | 6.9766 | 8.1444 | 7.2755 | 5.3826 | 7.3248 | 5.8596 | 6.6714 | 5.3044 | 10.2826 | 8.2626 | 10.1365 | 12.6641 | 14.8728 | 11.8154 | 15.0935 | 9.3742 | 8.8089 | 9.2193 | 8.2644
MOSR | 3.8483 | 3.6651 | 3.6163 | 3.3772 | 3.9215 | 4.0369 | 4.0679 | 4.4228 | 4.4914 | 3.8822 | 3.516 | 4.8077 | 3.906 | 3.6436 | 4.2301 | 4.0602 | 5.4989 | 4.409 | 4.5993 | 5.4174 | 5.1007 | 5.383 | 6.989 | 6.9011 | 6.3454 | 6.5107 | 6.2905 | 5.0104 | 5.2737 | 5.3915
Fig. 2. Each point is the SD of 30 measured steps in one 20× scan. The 30 scans retain their individual results for all three methods.
MOSR reduces Step SD in all 30 scans, with per-scan reductions of 36.9–64.6% relative to Raw. The within-scan summary is 8.431 nm for fixed Wiener and 4.865 nm for MOSR.
50×, 1 µm: 30 locked pairs
23.5%
5.938 nm Raw
4.542 nm MOSR
Reduction in Step SD
Chart: Step SD (nm)
| 50×, 1 µm
Raw | 5.938
Wiener | 4.7217
MOSR | 4.5417
Fig. 3. SD of the 30 locked 1 µm displacements at 50×. The same calibrated optical input supports Wiener and MOSR; the circle-center calculation is shared by all methods.
At 20×, MOSR lowers pooled within-scan step variation by 52.3% and improves all 30 repeated scans. At 50×, Step SD falls by 23.5% for the fixed 1 µm condition.
One reconstruction model supports both objectives with their calibrated optical inputs. Clearer boundaries and reduced displacement variation support MOSR for cross-magnification optical inspection.
[1] O. Ronneberger et al., U-Net, MICCAI, 2015, pp. 234–241.
[2] R. Zhang, Making convolutional networks shift-invariant again, ICML, 2019, pp. 7324–7334.
[3] T. F. Chan and L. A. Vese, Active contours without edges, IEEE Trans. Image Process., 10(2), 266–277, 2001.
[4] N. Chernov and C. Lesort, Least squares fitting of circles, J. Math. Imaging Vis., 23, 239–252, 2005.
Notes: Public model: MOSR. Source model identity: unchanged delivered R152 checkpoint.
Data: C:\Users\Rio\.codex\worktrees\impact-r152\GMM_2_Final\Experiments\R152_IMPACT2026Validation\results\separate_conditions\poster_data.json
20x: September 3 thesis cohort, 30 complete scans, 30 observed adjacent steps each. The line chart shows each scan SD. Summary is pooled within-scan SD, sqrt(mean(scan SD squared)); scans contain equal counts. This is not an average across acquisition conditions.
50x: July 2 original 1 micrometer cohort, 30 Raw-locked adjacent pairs. The bar chart is one sample SD per method from those 30 steps.
These are separate acquisition cohorts, not a matched causal isolation of magnification. All three methods share identical inputs and the same native normal-gradient geometric LSC reader within each condition.
The 0.5 micrometer 50x condition is retained in the complete audit and supplementary results; the submitted manuscript is unchanged.
All measured source numbers, reconstruction checks and citations are in the independent experiment workspace. The submitted manuscript supplies references [1]-[4].