## Slide 1
Methods
Conclusions
Measurement-Oriented Microscopy Reconstruction
for Cross-Magnification Micro-Displacement Inspection
Bibliography
Introduction
Results
Paper Code:
Yen Hsuan Chiang, Liang-Chia Chen*
Department of Mechanical Engineering, National Taiwan University
No. 1, Sec. 4, Roosevelt Rd., Taipei 10617, Taiwan. *Correspondence: lchen@ntu.edu.tw
Optical inspection estimates motion from image features. A sharper reconstructed edge is useful when its measured position remains stable. Optical blur, pixel sampling and reconstruction-induced boundary shifts can all change the reported displacement, especially when an inspection workflow switches objectives.
MOSR combines U-Net context [1], calibrated optical guidance and anti-aliased processing [2] to address these effects. Bounded correction retains the original image grid and limits geometry changes while improving boundary visibility. One fixed set of learned weights serves both objectives.
At a common nominal step of 1 µm, this study compares Raw, fixed Wiener and MOSR at 20× and 50×. Each condition contributes 30 measured single steps. The central outcome is the variation of these steps, linking reconstruction quality directly to displacement measurement.
Raw
MOSR
20×
2 µm
2 µm
50×
2 µm
2 µm
Fig. 1. Raw and MOSR images from the measured 20× and 50×, 1 µm datasets. Paired images use identical fields and grayscale ranges. Scale bars: 2 µm.
Data and reconstruction. At 20×, 31 frames form 30 consecutive steps; 50× uses the original 30 locked adjacent pairs. Both conditions employ a 1 µm nominal step. MOSR uses fixed weights and native-grid outputs, with 25 measured PSF kernels at 20× and one calibrated kernel at 50×. Fixed Wiener uses a regularization value of 0.0565.
Boundary and center. A coarse active contour [3] initializes each pillar boundary. Intensity profiles normal to the contour use 0.1-pixel spacing. Quadratic refinement locates each gradient maximum between samples. Geometric least-squares circle fitting [4] uses 64 refined points at 20× and 96 at 50× to estimate the center.
Displacement and statistics. Coordinate-wise medians of nine pillar-center changes define the array shift. Its component along the acquisition motion direction gives the measured step. Object-space conversion uses the scale in Table I. Each condition yields one sample SD from its 30 steps. Raw, Wiener and MOSR share the same images, features and center-calculation settings.
TABLE I. SEPARATE ACQUISITION CONDITIONS
Objective QA | Step (µm) | Measured steps | µm/pixel
20× | 1.0 | 30 | 0.1725
50× | 1.0 | 30 | 0.0690
Step SD uses all 30 displacements with denominator 29. Reduction (%) = 100 × (1 − SD(MOSR) / SD(Raw)). Each trace plots measured step minus the nominal 1 µm; each objective is summarized separately.
At a common 1 µm step, MOSR lowers displacement variation at both objectives. Results are calculated within each acquisition condition.
20×, 1 µm: 30 single steps
57.2%
9.506 nm Raw
4.068 nm MOSR
Reduction in Step SD
Chart: Step number
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Raw QA | -5.5006 | -5.9084 | 2.6995 | 4.1069 | 8.6555 | 2.3965 | -8.5655 | 0.50842 | 12.6509 | -5.7355 | 11.7848 | -26.0693 | 5.6064 | 10.0253 | -11.4786 | 20.235 | -8.7597 | -4.1751 | 10.5609 | -5.8781 | 4.0199 | 3.7729 | 4.2599 | 0.61785 | 7.1367 | 0.21894 | 9.489 | 11.022 | -10.847 | 9.7751
Wiener | 1.2162 | 2.7172 | -5.0176 | -4.7975 | -6.4943 | 8.0349 | 2.3345 | -0.051228 | -4.9412 | -2.722 | 5.5867 | 6.5633 | 2.9122 | 2.0684 | -1.4306 | 2.4327 | 17.393 | 9.439 | 7.8415 | -1.5315 | 2.7029 | -0.045107 | 12.3826 | -0.034985 | 0.03399 | -3.4392 | 11.2832 | -2.53 | 1.6233 | -5.1788
MOSR | -3.2279 | 2.711 | -2.7459 | -10.3681 | -3.9246 | 9.6615 | -1.8698 | -3.6535 | 2.0791 | 1.7639 | -0.20951 | 6.0874 | 0.2095 | 2.5216 | 4.6349 | 5.6154 | 0.2586 | 1.5373 | 1.8375 | 1.1707 | 1.5935 | 1.4932 | 5.0604 | -2.3249 | 8.5974 | 4.2866 | 0.36311 | 5.2617 | 4.0812 | 1.1231
Fig. 2. Thirty single-step displacements at 20×. The traces show measured step minus 1 µm, using the same axes and units as the 50× experiment.
At 20×, Step SD decreases from 9.506 nm for Raw to 5.794 nm for fixed Wiener and 4.068 nm for MOSR. MOSR reduces the variation by 57.2% relative to Raw.
50×, 1 µm: 30 single steps
23.5%
5.938 nm Raw
4.542 nm MOSR
Reduction in Step SD
Chart: Pair number
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Raw QA | 2.6864 | 9.1825 | 2.1385 | 19.0385 | 7.5524 | 8.2925 | 1.9384 | 0.16976 | 12.5169 | -7.5246 | 7.5557 | -6.0163 | 2.9378 | 4.85 | -1.9919 | 7.0065 | 10.3856 | 4.3704 | -4.6052 | 1.7295 | 6.64 | 2.9528 | 1.5653 | -1.8421 | 11.0021 | 7.2659 | -3.047 | 12.2517 | 1.9229 | 4.4843
Wiener | 0.77691 | 9.5928 | 0.73385 | 9.536 | 6.0438 | 8.9378 | 5.0386 | -2.235 | 8.7416 | -9.1882 | 4.837 | 2.3155 | 0.54845 | 0.92233 | 1.6564 | 7.4694 | 8.8941 | 1.794 | -7.0049 | 0.58665 | 6.6263 | 1.4774 | 1.9809 | 2.5077 | 3.5725 | 8.7267 | 3.9027 | 10.3701 | 7.7915 | 4.2995
MOSR | 0.68023 | 12.2429 | 1.8637 | 12.5647 | 5.4357 | 7.2084 | 4.6233 | -1.3055 | 7.424 | -7.9767 | 3.4927 | -1.5366 | 0.76876 | 0.78268 | -0.28674 | 4.945 | 9.1927 | 1.0557 | -5.1531 | -0.64555 | 6.0376 | 0.46536 | 4.1592 | 3.2941 | 2.3063 | 6.3042 | 2.7821 | 8.7737 | 3.9204 | 2.0864
Fig. 3. Thirty locked single-step displacements at 50×. Step SD is 5.938 nm for Raw, 4.722 nm for fixed Wiener and 4.542 nm for MOSR.
MOSR reduces the SD of 30 measured single steps by 57.2% at 20× and 23.5% at 50×, at a common nominal step of 1 µm. Each result retains its own acquisition condition.
One reconstruction model supports both objectives with calibrated optical inputs. Clearer boundaries and lower step variation support cross-magnification optical inspection.
[1] O. Ronneberger et al., U-Net, MICCAI, 2015, pp. 234–241.
[2] R. Zhang, Making convolutional networks shift-invariant again, ICML, 2019, pp. 7324–7334.
[3] T. F. Chan and L. A. Vese, Active contours without edges, IEEE Trans. Image Process., 10(2), 266–277, 2001.
[4] N. Chernov and C. Lesort, Least squares fitting of circles, J. Math. Imaging Vis., 23, 239–252, 2005.
Notes: Public model: MOSR. Source model identity: unchanged delivered R152 checkpoint.
Data: C:\Users\Rio\.codex\worktrees\impact-r152\GMM_2_Final\Experiments\R152_IMPACT2026Validation\results\separate_conditions\poster_data_30steps.json
20x: September 3 thesis run 003412_XY_0, acquisition run 7, the first evaluation run after the existing first-six-RAW-scan direction calibration split. Use all 31 frames and all 30 adjacent steps. This presentation selection follows the user request after full 30-scan auditing and is not a preregistered original-study choice.
50x: July 2 original 1 micrometer cohort, all 30 original Raw-locked adjacent pairs, ordered by acquisition index. They are selected pairs rather than a new 31-frame scan.
Each headline is 100*(1-SD(MOSR)/SD(Raw)), using sample SD of exactly 30 image-measured steps in that condition. There is no pooled scan statistic and no cross-objective average on this poster.
Curve values are observed image-measured displacements minus nominal 1000 nm, with identical axes in both plots. No per-method recentering, detrending, scaling, clipping or outlier removal is applied.
The complete 30-scan 20x study and 0.5 micrometer 50x result remain in the supplementary records. The submitted two-page manuscript is unchanged.
Both methods use the same normal-gradient geometric LSC family with the same settings within each condition. The submitted manuscript supplies references [1]-[4].