## Slide 1
Methods
Conclusions
Measurement-Oriented Microscopy Reconstruction
for Cross-Magnification Micro-Displacement Inspection
Bibliography
Introduction
Results
Paper Code:
Yen Hsuan Chiang, Liang-Chia Chen*
Department of Mechanical Engineering, National Taiwan University
No. 1, Sec. 4, Roosevelt Rd., Taipei 10617, Taiwan. *Correspondence: lchen@ntu.edu.tw
Optical inspection benefits from reconstruction when clearer boundaries produce more stable position measurements. Changing magnification changes object-space sampling and the pixel phase of the same motion. Single-step displacement dispersion therefore provides a direct test of measurement utility.
MOSR reconstructs each monochrome image on its original pixel grid. One fixed set of learned weights serves 20× and 50× objectives. Raw, fixed Wiener and MOSR use the same image pairs and pillar identities.
The primary outcome is one standard deviation from 30 locked single-step displacements in each condition.
Raw
MOSR
20×
2 µm
2 µm
50×
2 µm
2 µm
Fig. 1. Native-grid Raw and MOSR pillar fields at 20× and 50×. The common display range preserves direct visual comparison. Scale bars: 2 µm.
Reconstruction. MOSR combines U-Net context [1], anti-aliasing [2] and bounded PSF-guided correction. The learned weights remain fixed across objectives. Optical inputs comprise 25 measured kernels at 20× and one calibrated kernel at 50×.
Localization. Coarse contours [3] initialize geometric least-squares circle (LSC) centers [4]. The common reader refines each pillar with 720 uniform radial directions, 0.05-pixel profile sampling and three center updates.
Displacement. Each image pair yields nine independent horizontal center shifts. After conversion to object space, their absolute median defines one single-step displacement. The same procedure applies to all three image methods.
TABLE I. FIXED ACQUISITION CONDITIONS
Objective | Step (µm) | Pairs | µm/pixel
50× | 0.5 | 30 | 0.0690
50× | 1.0 | 30 | 0.0690
20× | 0.5 | 30 | 0.1725
20× | 1.0 | 30 | 0.1725
Within each condition, Step SD is the sample standard deviation of its 30 single-step displacements (denominator 29). The reported mean is the arithmetic mean of the four condition-level SDs.
Each condition contributes one Step SD from 30 measured single steps. MOSR yields the lowest Step SD in all four conditions, directly linking reconstruction to reduced step-to-step dispersion.
19.5%
8.84 nm Raw
7.11 nm MOSR
Reduction in the mean Step SD
Chart: Objective / nominal step (µm)
| 50× / 0.5 | 50× / 1.0 | 20× / 0.5 | 20× / 1.0
Raw | 5.2379 | 5.4301 | 14.1622 | 10.5278
Wiener | 4.907 | 4.5897 | 11.0768 | 11.0065
MOSR | 4.7227 | 4.5472 | 10.7543 | 8.4292
Fig. 2. Condition-level standard deviation of 30 single-step displacements. All methods use the same locked pairs and circle-center reader.
The reduction relative to Raw spans 9.8–24.1% across the four conditions. The largest reduction occurs at 20× with 0.5 µm steps, where Step SD decreases from 14.16 to 10.75 nm. MOSR also improves both 50× conditions.
TABLE II. STEP SD AND RELATIVE REDUCTION
Condition | Raw (nm) | Wiener (nm) | MOSR (nm) | Reduction vs Raw (%)
50× / 0.5 | 5.24 | 4.91 | 4.72 | 9.8
50× / 1.0 | 5.43 | 4.59 | 4.55 | 16.3
20× / 0.5 | 14.16 | 11.08 | 10.75 | 24.1
20× / 1.0 | 10.53 | 11.01 | 8.43 | 19.9
Mean | 8.84 | 7.90 | 7.11 | 19.5
Mean row: arithmetic mean of the four SDs. Reduction = (Raw − MOSR) / Raw.
MOSR lowers single-step dispersion in every tested condition, with reductions of 9.8–24.1% relative to Raw. The mean condition-level Step SD decreases by 19.5%.
One set of reconstruction weights supports both objectives. Clearer pillar boundaries and lower displacement dispersion establish MOSR as a measurement-oriented reconstruction approach for cross-magnification optical inspection.
[1] O. Ronneberger et al., U-Net, MICCAI, 2015, pp. 234–241.
[2] R. Zhang, Making convolutional networks shift-invariant again, ICML, 2019, pp. 7324–7334.
[3] T. F. Chan and L. A. Vese, Active contours without edges, IEEE Trans. Image Process., 10(2), 266–277, 2001.
[4] N. Chernov and C. Lesort, Least squares fitting of circles, J. Math. Imaging Vis., 23, 239–252, 2005.
Notes: Public model name: MOSR. Internal provenance maps to the fixed R152 delivery checkpoint in the experiment record.
User selected the original four-condition dataset with 30 single steps per condition; the separate 30-repeat-scan dataset is not part of this poster.
Source summary: C:\Users\Rio\.codex\worktrees\impact-r152\GMM_2_Final\Experiments\R152_IMPACT2026Validation\results\localization_improvement\improved_summary.csv
Source pair data: C:\Users\Rio\.codex\worktrees\impact-r152\GMM_2_Final\Experiments\R152_IMPACT2026Validation\results\localization_improvement\improved_pairs.csv
Primary quantity: one sample SD (ddof=1) of 30 image-measured single-step displacements per condition. The four SDs are averaged arithmetically. Percentage reduction of the mean uses the ratio of those means.
Images are the first locked 0.5 µm frame at each objective. Common crop and grayscale mapping; exact image paths remain in the source inspection manifest.
Chart style follows the submitted manuscript: Times New Roman chart text, monochrome patterns, explicit native-pixel scales, zero-baseline axes, and manuscript-style captions and table rules.
Template: C:/Users/Rio/OneDrive/學業/碩班/論文/投稿/IMPACT2026/4. Poster_Sample_2026 (2).pptx